## 12 bit A/D Converter

+9v

Syringe Driver (+ve)

Syringe Driver (-ve)

0v / GND

I = 80mA Max

Signal in

R2 = 4.3ohms

0v / GND

18 to 25

2

1

3

4

8

7

6

5

10F

4.7F

MAX187

Vdd

Ain

SHDN

Ref

GND

Dout

CS

SCLK

Parallel Port

2

9

11

5v

+9v

10F

78L05

9v dc Supply

## 5-10v Variable regulator using +5V 7805 regulator.

If the voltage at the "common" pin is raised above ground, this is added to the 5v regulator output, up to a limit about 2v below the input voltage – so a 12v supply can give 5-10v output.

The method shown is with a two-resistor potential divider on the output. Voltage across R1 is 5v, which produces a current of 5/R1 amps. This, plus the regulator's current from the "common" terminal (quiescent current Iq), flows through R2 producing a voltage across R2 of R2(5/R1 + Iq) volts. This voltage across R2 is added to the voltage across R1 (5v) to give the raised output.

The formula for the output voltage Vo is

Vo = 5(1 + R2/R1) + (Iq * R2)

Iq is typically 5mA (.005A), so if R1 is 240 ohms and R2 is 150 ohms,

## Vo = 5(1 + 150/240) + (.005 * 150) = 8.9v

For stability the divider current needs to be significantly higher than Iq, so R1 should be under 330 ohms.

# MAX187 Pinout

| 1 | VDD | Supply voltage, +5V ±5% |
| --- | --- | --- |
| 2 | AIN | Sampling analog input, 0V to VREF range |
| 3 | SHDN | Three-level shutdown input. Pulling SHDN low shuts the MAX187/MAX189 down to 10μA (max) supply current. Both MAX187 and MAX189 are fully operational with either SHDN high or floating. For the MAX187, pulling SHDN high enables the internal reference, and letting SHDN float disables the internal reference and allows for the use of an external reference. |
| 4 | REF | Reference voltage—sets analog voltage range and functions as a 4.096V output for the MAX187 with enabled internal reference. Bypass with 4.7μF if internal reference is used, and with 0.1μF if an external reference is applied. |
| 5 | GND | Analog and digital ground |
| 6 | DOUT | DOUT Serial data output. Data changes state at SCLK’s falling edge. |
| 7 | CS | Active-low chip select initiates conversions on the falling edge. When CS is high, DOUT is high impedance. |
| 8 | SCLK | Serial clock input. Clocks data out with rates up to 5MHz. |

## Programming Procedure

- Set CS pin low.

- Wait for Dout to go High (Data is ready)

- Set SCLK high then Low

- Read binary level on Dout. (First bit - MSB)

- Repeat 3 & 4 a further 11 times (bits 2 to 12).

For next conversion

- Set CS pin high

- Repeat steps 1 to 5 above

## ADC Input Circuit

12 Bit ADC = 4095 levels

Input range = 4v

Resolution = 4/4095 = 0.98mV/bit

## Graseby MS16/MS26 monitoring

0vv

+9vv

Syringe Driver (MS16 / MS26)

R1

Vi

Vi = Is * R1

If R1=1ohm

Is=20mA : Vi=0.02V (20mV)Digital : 20

Is=100Ma : Vi=0.1V (100mV)Digital : 102

Is=400mA: Vi=0.4V (400mV)Digital : 408

Is

Therefore the above circuit can be used to monitor the current flow through R1.  If the value of R1 is accurately known then the current will be:

Is = Vi * R1

The resolution of measurement will be better than 1mA.

Ain

ADC Output/Control Parallel Port Interface

9v PSU Input socket

+9v out (To Driver)

ON / OFF

Toggle Switch

Current Shunt

ON

0v (To Driver)

OFF